SITE- SPECIFIC Laser Safe Work Procedure:

EMERGENCY CONTACTS AND EQUIPMENT

A telephone is located along the north wall of the lab, outside the laser controlled area.

Emergency contact numbers:

Emergency (Security Services) :  555

Departmental Owner:

Laser Safety Contact:

Department Head:

LASER DESCRIPTION:

LASER SAFE WORK PROCEDURE:

Laser goggles are stored in_______. These goggles must be worn when the laser is in use by anyone within the laser controlled area. The correct goggles for alignment have an optical density of ___ at ___ nm; the goggles used for operation at higher power have an optical density of  ___ at ___ nm.

In the general Laser Safe Work Procedures (SWPs) described below, the Laser operator must insure that:

- Warning signage is posted on the entrance door and near the entrance to the laser controlled area

- The warning light is functioning when the laser controller is turned on

- A copy of the Laser SWPs is located on the computer desk with emergency contact information on front cover

- The telephone location on the north wall is know to all occupants

- All occupants in the laser controlled area during operation have appropriate laser eyeware and are listed as trained staff in Appendix B

- No extension cords are crossing walkways

- The “Panic Button” is located beside the laser during alignment and the computer during operation

- All reflective jewellery is removed by occupants

- The laser housing and beam stopping materials are in place

- All body parts are kept outside of the beam stopping materials during operation

- While operating, all occupants remain at the computer to avoid stray reflections

See Appendix A for a full risk hazard assessment.

APPENDIX A

SUMMARY OF HAZARDS: RISK ASSESSMENT VS. CONTROL MEASURES

LASER

| LASER DESCRIPTION | TYPE | CLASS |
| --- | --- | --- |
| U OF M # | WAVE LENGTH | MAXIMUM POWER |
| BUILDING | ROOM# | FACILITY |

RESPONSIBILITY FOR IMPLEMENTING CONTROL MEASURES

| POSITION | NAME | APPLICABILITY |
| --- | --- | --- |
| DEPARTMENT HEAD |  | NO |
| DEPARTMENTAL OWNER |  | YES |
| LASER OPERATORS | SEE APPENDIX B | YES |
| OTHER OCCUPANTS | SEE APPENDIX B | NO |

HAZARDS

| HAZARD (CONTROLLED AREA) | RISKS | CONTROL MEASURES |
| --- | --- | --- |
| DIRECT LASER RADIATION, DIFFUSE AND SPECULAR REFLECTIONS FROM CLASS 4 LASER | LOW TO MEDIUM POTENTIAL FOR EYE AND/OR SKIN DAMAGE RISK TO EPILEPTICS | LASER CONTROLLED AREA AS PER ANSI Z136.1-2007: DEFINED AS THE FLOOR SPACE WITHIN room EAST OF THE LASER CONTROLLED AREA BARRIER (NOMINAL HAZARD ZONE) WARNING SIGNAGE: WARNING LIGHTS: AMBER STROBE LIGHT LOCATED ON THE OUTSIDE OF LASER CONTROLLED AREA BARRIER ACTIVATES WHEN LASER CONTROL UNIT IS TURNED ON - ENGINEERED. ENTRY CONTROL: NON-DEFEATABLE, INTERLOCKED WITH LASER ENGINEERED. EMERGENCY EXITING: THROUGH LASER CONTROLLED AREA BARRIER - CURTAIN ON ROLLER TRACK INTERLOCKED WITH LASER- ENGINEERED. EMERGENCY LASER SHUTDOWN: MOVEABLE, RED MUSHROOM SHAPED “PANIC BUTTON” - ENGINEERED. PROTECTIVE EYEWARE: ALL PERSONS IN LASER CONTROLLED AREA DURING OPERATION HAVE APPROPRIATE EYEWARE - PROCEDURAL MAXIMUM OCCUPANCY 2 GOGGLE/SPECTACLE SPECIFICATIONS, 1) GLENDALE 2) GLENDALE 190-520nm OD+9 Same as 1) 520-532nm OD+7 710-750nm OD+3 750-850nm OD+5 850-1080nm OD+7 5000-11000nm OD+7 LASER KEY CONTROL: LASER KEY KEPT IN CABINET A -ACCESSIBLE TO AUTHORIZED LASER OPERATORS ONLY (SEE APPENDIX B FOR LIST) - PROCEDURAL ACCESS CONTROL: ONLY TRAINED STAFF ARE PERMITTED IN THE LASER CONTROLLED AREA DURING OPERATION PROCEDURAL TRAINED STAFF-EHSO LASER SAFETY INFORMATION PPT COMPLETED (SEE APPENDIX B FOR LIST) FOR PROTECTION OF ANCILLARY STAFF (UNTRAINED) SIGNAGE, KEY CONTROL, STROBE LIGHT, ENTRY INTERLOCK SAFE WORK PROCEDURES(SWPS): COPY LOCATED IN BINDER ON PIV CONTROL COMPUTER DESK WITH EMERGENCY CONTACT INFORMATION UNATTENDED OPERATION IS PROHIBITED |

| HAZARD (LASER - NON BEAM) | RISKS | CONTROL MEASURES |
| --- | --- | --- |
| ELECTRICAL / FLOOD | 115V 10 AMP 1PH POWER SUPPLY TO LASER WITH INTERNAL WATER SUPPLY FOR COOLING- LOW POTENTIAL FOR WATER LEAKAGE TO CONTACT HIGH POWER | CSA APPROVED GROUNDED ELECTRICAL SERVICE ENGINEERED. LASER CONTROLLER PLACED ON WOOD BLOCKS ABOVE FLOOR LEVEL - PROCEDURAL EXTENSION CORDS SURGE PROTECTED FLOOR LEVEL DRAIN IN VICINTY OF LASER AND WATER TANKS - ENGINEERED. |
| FIRE: | MOST STRUCTURES IN THE LASER CONTROLLED AREA ARE METALLIC WITH SOME FIBERGLASS, PLASTICS AND WOOD. 120 mJ PULSED LASER HAS LOW POTENTIAL TO START A FIRE | SPRINKLERS WITHIN LASER CONTROLLED AREA (HEAT SENSING) - ENGINEERED. CONTROLLED AREA BARRIER AT LASER LEVEL CONSTRUCTED FROM NON-FLAMMABLE SILICON RUBBER IMPREGNATED FIBERGLASS WELDING BLANKET, GALVANIZED STEEL STRUCTURAL MEMBERS. UPPER PANELS ARE CELTEC PVC FOAM –CAN NOT SUSTAIN A FLAME - ENGINEERED. ALL BUILDING WALLS SHEETROCK - ENGINEERED. NO FIRE EXTINGUSHER FOR LOCAL CONTROL LOCATED IN E1-547, NOT REQUIRED |
| HAZARD (APPLICATION) | RISKS | CONTROL MEASURES |
| ANALYZING WATER JET INTO TANK USING PIV (PARTICLE IMAGE VELOCIMETRY) | NOTE: DAMAGE TO CCD FROM INTENSE LIGHT | USE APPROPRIATE FILTERS |
| FLOOD CREATES: SHOCK /SLIPPING | HIGH POTENTIAL FOR WATER LEAKAGE FROM ACRYLIC TANK, HOSES, PUMP ONTO FLOOR MEDIUM POTENTIAL FOR SLIPPING LOW POTENTIAL FOR SHOCK | CSA APPROVED GROUNDED ELECTRICAL SERVICE ENGINEERED. LASER CONTROLLER PLACED ON WOOD BLOCKS ABOVE FLOOR LEVEL - PROCEDURAL EXTENSION CORDS SURGE PROTECTED FLOOR LEVEL DRAIN IN VICINTY OF WATER TANKS ENGINEERED. SPILLS CLEANED UP PROMPTLY - PROCEDURAL |
| NOISE | THE NOISE LEVEL IS CREATED BY COMPUTERS AND SMALL ELECTRIC PUMPS – LOW POTENTIAL FOR HEARING DAMAGE | USE OF HEARING PROTECTION DURING TESTING IS NOT REQUIRED |
| AIR QUALITY (FOR WATER USE ONLY) | ACTIVITIES ASSOCIATED WITH THIS TEST DO NOT CREATE A CHANGE IN THE AIR QUALITY | ADDITIONAL VENTILATION, RESPIRATOR OR MASK USE IS NOT REQUIRED |
| TRIPPING | EXTENSION CORDS , HOSES CROSSING WALKWAYS- MEDIUM POTENTIAL FOR TRIPPING | EXTENSION CORDS(INCLUDES SURGE PROTECTION) FROM ELECTRICAL SERVICE TO LASER CONTROL AND PUMP ALONG WITH DATA CABLES ARE SUSPENDED OVERHEAD ACROSS WALKWAYS OR COILED AWAY FROM WALKWAY- UNDER TABLES PROCEDURAL |

APPENDIX B

LASER OPERATORS

| NAME | TRAINED BY | DATE |
| --- | --- | --- |

EHSO LASER SAFETY INFORMATION PPT COMPLETED

| NAME | CERTIFICATE SIGNED | DATE |
| --- | --- | --- |